## Sensors
| FORMAT KEY: | Good aspect | Problematic aspect / device | Caution
USER INTERACE
| Concept | Hardware | Qty @ Cost | Ease of Use | Scalability | Lifetime | Comments
| "Turn & Click" / Potentiometer | PEC11R-4225F-S0024 | 1 @ $1.53 | Requires menu navigation for all settings (SLOW) | Single control can be used for anything. | Rotation: >30k cycles Button: >20k cycles | Elegant!
| Buttons | Pick based on final form factor. Something like: MX1A-11NN | At least 5-6 @ ~$1.00 ea. (SMD are less) | Intuitive up/down buttons for various parameters. | More features require more buttons. | >50,000,000 cycles
| Capacitive Touch Sensors | Sensor pads can be constructed with conductive tape / foil | None (implemented in SW) | Intuitive up/down (capacitive) buttons for various parameters | More features require more (capacitive) buttons… but this is not expensive. | Ideally infinite | Need to carefully test to ensure sensitivity with/without PPE, etc.
PRESSURE SENSOR
| Component | Form Factor | Cost | Type | Range & Accuracy | Digital Interface | Comments | Link
| BME280 | 2.5x2.5mm surface mount package w/ metal lid | $6.55 sensor $20 SparkFun breakout | Absolute | 300 ~ 1100 hPa +/- 1.0 hPa (0.12 hPa typ.) | I2C and SPI | Includes humidity + temp. sensors | https://ae-bst.resource.bosch.com/media/_tech/media/datasheets/BST-BME280-DS002.pdf
| BMP180 |  | Obsolete / Out of stock
| MS5803-14BA |  | $60 SparkFun breakout | Absolute | Error = +/- 20 hPA at best |  |  | https://www.sparkfun.com/products/12909
| MPXV7002 | 8-SOP, 8mm tall w/ 2 male ports | $16.90 sensor | Differential | Range: +/- 20 hPA +/- 2.5 hPa (1.0 hPa typ.) | Analog Output | Temperature Compensated | https://www.digikey.com/products/en?mpart=MPXV7002DP&v=568
| MPXV5010DP | 8-SOP, 10mm tall w/ 2 male ports | $17.70 sensor | Differential | Range: 0 ~ 100 hPA Error: 5 hPa | Analog Output | Temperature Compensated | https://uk.rs-online.com/web/p/differential-pressure-sensor-ics/7191184
| ADP5100 |  |  | Differential | Range: +/- 1000 hPa Error: 13 hPa |  |  | https://uk.rs-online.com/web/p/gauge-pressure-sensor-ics/8368931
MS4525 Series | MS45254525-DS5A001DP |  | $39.31 sensor | Differential | Range: 0 ~ 68.9 hPa 0.17 hPa |  |  | https://www.te.com/usa-en/product-CAT-BLPS0041.html
| MPXC2011TD1 | SIL4 | 5.82 | Differential | Range: 0-100 hPa | Differential Analog | Approved for medical (vents in particular) Temperature Compensated | https://www.findchips.com/search/mpxc2011dt1
Pressure Sensor Notes: UK requested: Normal pressure of 35 cmH2O, adjustable in increments of 5 cmH2O up to a potential maximum of 70 cm H2O (with PEEP in the range 5-20 cmH2O) 1 cmH2O = 0.98 hPa Atmospheric pressure is 1,013 hPA
FLOW RATE SENSOR
| S-type Pitot tube using two pressure sensors? | https://www.sciencedirect.com/science/article/abs/pii/S0955598614001551

## Readouts
VISUAL READOUT
| Component | Type / Size | Cost | Interface | Notes
| Sparkfun Nokia-5110 Breakout | 84x48 pix LCD | 9.95 | SPI | -Backlight -Arduino Example: https://playground.arduino.cc/Code/PCD8544/ | https://www.digikey.com/product-detail/en/sparkfun-electronics/LCD-10168/1568-1349-ND/5764484
| COG-C144MVGI-08 | 128x128 pix LCD | 7.05 | I2C | -Backlight -**Note: Has a specialized gold-finger connector -Have not found a datasheet
| Adafruit I2C OLED Display | 128x32 pix OLED | 9.95 (old) or 12.50 (new) | I2C | -LED Display (doesn't need backlight) -"New" version has QWIC connector + auto-reset -Adafruit tutorials (compatible w/ Arduino)
ALARM
| Component | Typ. Loudness | Cost | Interface/Control | Notes
| Sparkfun Large Piezo Alarm | 75 dB @ 3V*5mA = 15mW 100dB @ 18V*40mA = 720 mW | 2.95 | DC Voltage |  | https://www.sparkfun.com/products/13940
| Sparkfun PCB mount speaker | 80 dB @ 100 mW | 1.95 | Audio Signal (8ohms)
| PKM22EPPH4012-B0 | 85dB @ 1.5V (power??) | 0.78 | DC Voltage |  | https://www.digikey.com/product-detail/en/murata-electronics/PKM22EPPH4012-B0/490-7032-ND/2595278